# RHTemp1000IS Intrinsic Safety User Notice

## Intrinsic Safety Approval:

The RHTemp1000IS has been certified by FM Approvals as Intrinsically Safe (IS) for use in Class I, Division 1, groups A, B, C, D, and Nonincendive (NI) for use in Class I, Division 2, groups A, B, C, D Hazardous (Classified) Locations. The rating listed in the Factory Mutual approval guide is as follows:

RHTemp1000is.  Temperature and Humidity Recorder.

IS / I / 1 / ABCD T4A Ta=80°C; NI / I / 2 / ABCD / T4A Ta=80°C

These are the only safety ratings relevant to the use of this product. Use of this product in hazardous environments not specifically covered by this rating, is prohibited, unless the user takes the appropriate steps to ensure the safety of the product and assumes full responsibility for its safe use. Refer to the reference sections at the end of this document for further information on approval standards and environments.

## Conditions of Use:

The following conditions must be satisfied to maintain the IS rating of the RHTemp1000IS:

- When used in hazardous locations, the RHTemp1000IS is to be installed prior to the location becoming hazardous, and removed only after the area is no longer hazardous.

- The maximum allowed ambient temperature for the RHTemp1000IS (under any circumstances) is 80°C. The minimum rated operating temperature is -20°C.

- The RHTemp1000IS approved for use only with the Tadiran TL-2150/S battery. Replacement with any other battery will void the safety rating.

- Batteries are user replaceable, but are to be removed or replaced only in locations known to be non-hazardous.

- Tampering or replacement of non-factory components may adversely affect the safe use of the product, and prohibited. Except for replacement of the battery, the user may not service the RHTemp1000IS. MadgeTech, Inc. or an authorized representative must perform all other service to the product.

DOC-1271036-IS

- The RHTemp1000IS enclosure does not carry a NEMA rating, as there are openings in the enclosure for measuring humidity. For this reason, the product is only suitable for use in indoor locations.

## Required Labeling:

The following label must be engraved to the enclosure of the RHTemp1000IS. It contains critical information for the safe use of the product.

## Reference Standards:

The RHTemp1000IS complies with the following standards:

| Standard | Date | Title |
| --- | --- | --- |
| FM Class 3600 | 2011 | Electrical Equipment for Use in Hazardous (Classified) Locations, General Requirements |
| FM Class 3610 | 2010 | Intrinsically Safe Apparatus and Associated Apparatus for Use in Class I, II and III, Division 1 and Class I, Zone 0 and 1 Hazardous (Classified) Locations |
| FM Class 3611 | 2004 | Non-incendive Electrical Equipment for Use in Class I and II, Division 2 and Class III, Divisions 1 and 2 Hazardous (Classified) Locations |
| FM Class 3810 Including Supplement #1 | 2005 1995 | Electrical Equipment for Measurement, Control, and Laboratory Use |

## Protection and Environment Reference:

The environmental rating is per ANSI/NFPA 70 National Electric Code® (NEC®) Article 500. The following is information excerpted from FM Approvals reference documents.

Protection Concepts

| Type of Protection | Code | Permitted Use | Standard |
| --- | --- | --- | --- |
| Intrinsic Safety | (IS) | Class I, Division 1 | FM Class 3610 |
| Non-Incendive | (NI) | Class I, Division 2 | FM Class 3611 |

Apparatus Grouping Per NEC® 500:

| Class | Group | Typical Gas |
| --- | --- | --- |
| I | A | Acetylene |
|  | B | Hydrogen |
|  | C | Ethylene |
|  | D | Propane |

Area Classification Per NEC® 500:

| Division | Description |
| --- | --- |
| 1 | Flammable Material Present Continuously |
|  | Flammable Material Present Intermittently |
| 2 | Flammable Material Present Abnormally |

Temperature Class Per NEC® 500:

| Temperature Class | Maximum Surface Temperature (of any component under fault conditions) |
| --- | --- |
| T4A | 120 °C (with maximum 80 °C ambient) |

Note: the T4A rating indicates the maximum surface temperature potentially encountered in a fault condition. This is not the allowed operating temperature. This temperature rating limits the maximum ambient temperature of 80°C.

6 Warner Road, Warner, NH 03278

Phone 603.456.2011 ( Fax 603.456.2012

www.madgetech.com ( support@madgetech.com